## Sheet1
TID (unique) | KEYWORD(unique) | TAG(Separate with comma) | CATEGORY | PID(Separate with comma if more than one) | QUESTION | ANSWER | Documentation link | YouTube solution link
GVR81 | Before Power On | power supply | power | VENR860 | procedures to be done before powering on the system ? | 1. Connect the Inspiratory Safety Guard (ISG) to the inspiratory port. 2. Connect the Exhalation Valve Assembly (EVA) to the exhalation port. 3. Familiarize yourself with the hard keys on the display. a. 2-minute audio pause b. Increase O2 c. Snapshot d. Lock/Unlock e. Home f. Trim Knob 4. Attach the breathing circuit and all accessories that will be used to ventilate the patient. | https://drive.google.com/file/d/1Gub9cUvd1MbbtEV6c5c1B4PTVqf510ed/view?usp=sharing
GVR82 | Power on | Turn on | power | VENR860 | How to turn on the system ? | 1. Connect the air and oxygen supply hoses to the gas supply source. 2. Plug power cord into an approved electrical outlet. 3. Press the power switch to turn on the ventilator. | https://drive.google.com/file/d/1vhVgeBiqC-KXBhWBB-tWE4tHDtspOb1r/view?usp=sharing
GVR83 | Patient setup | patient care | interface | VENR860 | How to setup the system for new patient ? | 1. Select NEW PATIENT. 2. Select Adult, Pediatric or Neonatal patient type. .3. Select Patient ID (Identification), if necessary, and enter 10 characters and select Confirm. 4. Select Gender (male or female). 5. Select Height. 6. Select Weight. 7. Select Tube Type (Endotrach, Trach, or ---. 8. Select Tube Diameter. 9. Verify and confirm settings by pressing Confirm. | https://drive.google.com/file/d/1aCSe_Rg8cx_0lJegShdeOev22xT3YM5z/view?usp=sharing
GVR84 | System Check | system check,system troubleshoot,circuit check up | Interface | VENR860 | How to perform system check ?
GVR85 | RUNNING_A_SYSTEM_CHECK | Run, troubleshooting | Interface | VENR860 | How to check the system in running | 1. From Standy, select system check. 2. Attach the breathing circuit and all accessories that will be used to ventilate the patient. 3. Occlude the patient wye using the occlusion port. 4. Select start 5. Select the information icon to see the system check details menu | https://drive.google.com/file/d/10lDZamRT-fhPTgs4zdNGXqwMxfH_ChTu/view?usp=drivesdk
GVR86 | SETTING_A_VENTILATION_MODE | setup, setting, ventilation | Interface | VENR860 | How to setting the ventilation mode | 1. Select current mode. 2. Select the desired ventilation mode. 3. Select Assist Control, Leak Comp or Trigger Comp, if desired. 4. Set the desired settings for the Ventilation mode and Confirm. | https://drive.google.com/file/d/10n2XhqFRIIcx9Cf18ncoRuz1zEP59ZKr/view?usp=drivesdk
GVR87 | PATIENT_VENTILATION | Ventilation, start | Interface | VENR860 | How to start the patient ventilation | 1. From Standby, select Start Ventilation. 2. After ventilation has started, connect the breathing circuit to the patient | https://drive.google.com/file/d/10p0H_eUxR5Gu36jUv9hHHfNRcUyY9X-m/view?usp=drivesdk
GVR88 | ALARM_PANEL | Alarm, indicator, panel | Indicator | VENR860 | How to use the Alarm panel | 1. Audio pause. 2. Alarm Status. 3. Alarm setup. 4. Audio Pause Timer. 5. Active Alarm count | https://drive.google.com/file/d/10r-ICQLMuIZTAkhXB9vk7nQ3H25qTCGh/view?usp=drivesdk
GVR89 | ALARM_SETUP | Setup, Alarm, settings | Indicator | VENR860 | How to set the Alarm | 1. Apnea Time: Set apnea time up to 60 seconds. 2. Alarm Light Brightness: Set the brightness level of the alarm light. The range is 1 (Low) to 5 (High). 3. Alarm Volume: Set the alarm volume of the alarm tone. The range is 1 (Low) to 5 (High). 4. High Alert Audio: Set the delay time in which a high-priority alarm must be resolved before the audio tone pitch and volume increases to the maximum level. The range is 0 to 30 seconds or Off. A general message shows when High Alert Audio is set to Off. 5. Alarm Limits: Select the check box to show alarm limits adjacent to the measured data in the patient data workspace. The alarm limit always shows when an alarm occurs for the measured data, even if it is set to Off. 6. Auto Limits: Select to set auto alarm limits based on current measured data for: a. MVexp-Low and High b. VTexp-Low and High c. Rate-Low and High d. EtCO2-Low and High e. PEEPe-Low and High | https://drive.google.com/file/d/10um-GV2sIbFX0vncHXUwqLv-E65nl8JS/view?usp=drivesdk
GLOGIQ91 | POWERING_THE_SYSTEM | power supply,power,system | power | LOGIQ9 | Procedures to be done before powering on the system ? | 1. Ensure that the wall outlet is of the appropriate type. 2. Ensure that the power switch is turned off. 3. Unwrap the power cable. Make sure to allow sufficient slack in the cable so that the plug is not pulled out of the wall if the system is moved slightly. 4. Attach the power plug to the system and secure it in place by using the retaining clamp
GLOGIQ92 | PASSWORD_PROTECTION | Password,Protection | Power | LOGIQ9 | How to set Password to the system? | 1. Operator: Select the Operator. 2. Password: Enter Operator’s password (optional). 3. Select type of Logon or Cancel. • Emergency: Data stored only for the duration of the current examination. • OK: Standard logon • Cancel: Cancel logon If the correct ID and Password are entered, the system displays the Patient screen. Logoff To logoff, press the Power On/Off switch momentarily and a SYSTEM-EXIT window appears.
GLOGIQ93 | STANDBY_MODE | Power,Time | Power | LOGIQ9 | How to Standby Mode? | 1. Press the Power On/Off switch at the front of the system once. 2. The System-Exit window displays. 3. Using the Trackball, select Standby. It takes approximately 20 seconds to go into Standby Mode. 4. Transport the system as needed and plug it back in. See ‘Moving the System’ for more information. 5. Press and immediately release the Power On/Off switch to come out of Standby Mode. This takes about 45 seconds.
GLOGIQ94 | POWER_OFF | Power | Power | LOGIQ9 | How to turn off the System | 1. Press the Power On/Off switch at the front of the system once. 2. The System-Exit window is displayed. 3. Using the Trackball, select Shutdown. The shutdown process takes a few seconds and is completed when the Power On/Off switch illumination is turned amber. 4. Disconnect the probes. Clean or disinfect all probes as necessary. Store them in their shipping cases to avoid damage.
GLOGIQ95 | CIRCUIT_BREAKER | Power | Power | LOGIQ9 | How to manage Power overloads? | If a power overload occurs: 1. Turn off all peripheral devices. 2. Reactivate the Circuit Breaker switch. The Circuit Breaker switch should stay in the On position; DO NOT hold the switch in the On position. If the Circuit Breaker switch remains On, follow the Power On procedure previously described. NOTE: If the Circuit Breaker switch does not remain in the On position or trips again: 1. Disconnect the Power Cable. 2. Call Service immediately. DO NOT attempt to use the system.
GLOGIQ96 | TRANSPORTING_THE_SYSTEM | Transpoting,Positioning | Transport | LOGIQ9 | Precautions taken before Transporting or Moving the System? | 1. Press the Power On/Off switch to power off. See ‘Power Off’ for more information. 2. Unplug the power cord. 3. All cables from off-board peripheral devices (IIE camera, external printer, VCR, etc.) must be disconnected from the console. 4. Disconnect the footswitch from the console. 5. Ensure that no loose items are left on the console. 6. Wind the power cable around the cable hook below the rear handle or wrap the cord in a bundle and store it behind the rear storage panel. 7. Connect all probes to be used while off site. Ensure that probe cables are out of the way from the wheels and not protruding beyond the console. NOTE: If more than four (4) probes are intended to be used, store the additional probes securely. 8. Store all other probes in their original cases or in soft cloth or foam to prevent damage. 9. Store sufficient gel and other essential accessories in the provided space. 10. Adjust the monitor and control panel to its lowest position possible by using the release lever located underneath the control panel. Make sure the keyboard is locked in place. 11. Unlock the wheels
GLOGIQ97 | PROBE_CONNECTION | Probe,Connection | Interface | LOGIQ9 | How to Connect the Probe safely? | 1. Place the probe's carrying case on a stable surface and open the case. 2. Carefully remove the probe and unwrap the probe cord. 3. DO NOT allow the probe head to hang free. Impact to the probe head could result in irreparable damage. Use the integrated cable management hook to wrap the cord. 4. Turn the connector locking handle counterclockwise. 5. Align the connector with the probe port and carefully push into place. 6. Turn the connector locking handle clockwise to secure the probe connector. 7. Carefully position the probe cord so it is free to move and is not resting on the floor.
GLOGIQ98 | PROBE_CABLE | Probe,Cable | Interface | LOGIQ9 | How to handle the Probe cable? | • Keep free from wheels • Do not bend the cable acutely • Avoid crossing cables between probes.
GLOGIQ99 | PROBE_ACTIVATION | Probe,Activation,Deactivation | interface | LOGIQ9 | How to manage probe? | To activate the probe, select the appropriate probe from the probe indicators on the Touch Panel. 1. Probe Indicators The probe's default settings for the mode and selected exam are used automatically. Deactivating the Probe When deactivating the probe, the probe is automatically placed in standby mode. Disconnecting the Probe Probes can be disconnected at any time. However, the probe should not be active when disconnecting the probe. • Ensure the probe is deactivated. Deactivate by selecting another probe or pressing Freeze. • Move the probe locking handle counterclockwise. • Pull the probe and connector straight out of the probe port. • Carefully slide the probe and connector away from the probe port and around the right side of the keyboard. • Ensure the cable is free. • Be sure that the probe head is clean before placing the probe in its storage box or a wall hanging unit. Freeze Deactivate the probe: 1. Press the Freeze key. 2. Gently wipe the excess gel from the face of the probe. 3. Carefully slide the probe around the right side of the keyboard, toward the probe holder. Ensure that the probe is placed gently in the probe holder.
GLOGIQ910 | CONTROL_PANEL_MAP | Control Panel | Control | LOGIQ9 | How to use Control panel? | 1. Touch panel and controls 2. Touch Panel Brightness Control 3. Video 4. Audio On/Off and Volume 5. TGC 6. Reverse 7. Additional Feature Keys 8. Keyboard 9. Mode/Gain Keys 10. Imaging/Measurement Keys 11. Depth 12. Imaging Feature Keys 13. Freeze and Print Keys 14. Probe and Cord Holder 15. Gel Holder
GLOGIQ911 | EXAM_FUNCTION_CONTROLS | Function,controls | Control | LOGIQ9 | How to manage exam function? | 1. Patient: Enters Patient screen 2. Scan: Enters scanning mode screen 3. Reports: Activates default report and Touch Panel of report choices. 4. End Exam: Activates Image Management and Touch Panel with end of exam options. 5. Utility: Activates system configuration menus. 6. Preset: Selects the application to use. 7. Probe Indicator: Indicates and selects the probes.
GLOGIQ912 | PATIENT_SEARCH | patient care | Exam | LOGIQ9 | How to search for an exsiting Patient? | 1. Press Patient to display the Patient Screen. The Patient Search Screen is displayed. 2. Select the search key and Location (From). Enter the search string. 3. An appropriate patient is displayed in the patient list. 4. Select the patient and press Review. or Click ‘+’ mark on the left side of Patient ID, then all exams of the patient is displayed in the patient list. Select the required exam and press Active Images. The image list of the exam is displayed. 5. The selected patient exam screen is displayed.
GLOGIQ913 | IMAGE_REVIEWING | Image | Interface | LOGIQ9 | How to review the Image? | 1. Select the exam which includes the image to review. 2. Press Active Images. 3. Select the image and press Review. 4. The image is displayed.
GLOGIQ914 | INSPECTING_THE_SYSTEM | System Care,Maintenance | Maintenance | LOGIQ9 | How to Examine the System? | • Connectors on cables for any mechanical defects. • Entire length of electrical and power cables for cuts or abrasions. • Equipment for loose or missing hardware. • Control panel and keyboard for defects. • Casters for proper locking operation. To avoid electrical shock hazard, do not remove panels or covers from console. This servicing must be performed by qualified service personnel. Failure to do so could cause serious injury.
GLOGIQ915 | CLEANING_AIR_FILTER | Air filter,Filter | Maintenance | LOGIQ9 | How to Clean the Air Filter? | Cleaning To clean the filter: 1. Be sure to shake the filter in an area away from the system. 2. Wash the filter in a mild soapy solution, rinse and air dry or dry with a cloth. NOTE: Allow wet filter to dry thoroughly before installing. Installing 1. Slide the filter back into the system. 2. Secure the filter with the Velcro tab.
GLOGIQ916 | SCANNING_A_PATIENT | Patient,Scanning | Exam | LOGIQ9 | How to Scan a New patient? | 1. Press Patient. 2. Press New Patient on the Patient menu. 3. A pop-up menu appears. Specify whether you want to store images permanently, delete images, or go to active images. 4. Choose the exam category. 5. Verify the dataflow. 6. Fill in patient information. 7. Press Register. Enter Past OB Exam information, if desired. 8. Select the probe from the Touch Panel to start scanning (or select Exit, Escape, Scan, or Freeze). 9. Perform the exam. 10. When you have completed the study, press End Exam, End Current Patient, then Store All Images Permanently.
GLOGIQ917 | PATIENT_LIST | Patient,List | Exam | LOGIQ9 | How to enter a Patient List? | 1. Press Patient to display Patient Screen. 2. Press New Patient to erase the current patient data. 3. Enter the Patient ID. 4. Enter the patient and exam information. 5. Press Register. 6. Repeat above steps as required. Select the patient from the Patient List and press New Exam to start an exam.
GLOGIQ918 | ERASE_EXISTING_PATIENT | Patient,Delete | Exam | LOGIQ9 | How to delete the Existing Patient? | 1. Search and select the patient in the patient list. 2. Press Delete. 3. The confirmation dialog box is displayed. 4. Press OK to delete or Cancel.
GLOGIQ919 | ERASE_EXISTING_EXAM | Exam,Delete | Exam | LOGIQ9 | How to Delete the Existing Exam List? | 1. Search and select the patient in the patient list. 2. Press Review. 3. The patient exam screen is displayed. Select the exam which you want to delete. 4. The confirmation dialog box is displayed. 5. Press OK to delete or Cancel.
GLOGIQ920 | ERASE_EXISTING_IMAGE | Image,Delete | Exam | LOGIQ9 | How to Delete the Existing Image list? | 1. Search and select the patient in the patient list. 2. Press Review. 3. The patient exam screen is displayed. Select the exam which include the image to delete. 4. Press Active Images to display the image list. 5. Select the image to delete and press Delete. 6. The confirmation dialog box is displayed. 7. Press Yes to delete or No to cancel.
GLOGIQ921 | DEAPTH_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust Depth? | Each adjustment cycles you to the next Depth setting. Imaging and display parameters adjust automatically. To increase/decrease, adjust Depth.
GLOGIQ922 | GAIN_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust Gain? | Gain values vary depending on the probe; they are not associated with a particular position of the knob. NOTE: TruAccess allows you to change the gain on a frozen image. B-Mode gain is independent of M-Mode and Doppler and Color Flow Gain. Changing the gain while in M-Mode does not affect the B-Mode image gain. To decrease/increase, rotate Gain.
GLOGIQ923 | FOCUS_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust Focus? | To increase/decrease the number of focal zones, adjust Focus Number. NOTE: Focus Position and Focus Number share the same control. Press the control to toggle between Focus Position and Focus Number. To move the focal zone to the near/far field, adjust Focus Position. You can reverse how the Focus control functions on the Utility -> System -> General page.
GLOGIQ924 | B_FLOW_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Adjust B flow? | To activate/deactivate B Flow, press B Flow. Doppler Mode is available while in B Flow; however, M-Mode and Color Flow/PDI Modes are not available.
GLOGIQ925 | SRI_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Ajust SRI? | To activate SRI, adjust the SRI Touch Panel.
GLOGIQ926 | TIC_B-MODE | Adjust | Optimizing the image | LOGIQ9 | How to Analyse TIC? | 1. Scan the patient after injecting the contrast agent. 2. Watch the agent flow through the anatomy of interest. 3. When the desired contrast effect has been visualized, freeze the image and select a range of images for analysis. 4. Position an ROI (region of interest) on one of those images where the contrast effect is visible. 5. The system then calculates the mean pixel intensity within that ROI for all frames in the user designated loop and plots the resulting data as a function of time.
GLOGIQ927 | ACTIVATING_TIC | Activation,Adjust | optimizing the image | LOGIQ9 | How to Activate TIC? | 1. Scan the patient in Contrast mode or select a desired cine loop from the stored images. 2. Press Freeze. 3. TIC Analysis displays on the Cine Touch Panel. 4. Select TIC Analysis. The TIC Analysis screen and Touch Panel displays. To toggle the trackball function between TIC and Scroll, press Scan Area.
GLOGIQ928 | EXAMINATION_M-MODE | Exam, | optimizing the image | LOGIQ9 | How to Examine using M-mode? | 1. Get a good B-Mode image. Survey the anatomy and place the area of interest near the center of the B-Mode image. 2. Press M/D Cursor. 3. Trackball to position the mode cursor over the area that you want to display in M-Mode. 4. Press M-Mode. 5. Adjust the Sweep Speed, TGC, Gain, Power Output, and Focus Position, as needed. 6. Press Freeze to stop the M trace. 7. Record the trace to disk or to the hard copy device. 8. Press Freeze to continue imaging. 9. To exit, press M-Mode.
GLOGIQ929 | EXAMINATION_DOPPLERMODE | Exam | optimizing the image | LOGIQ9 | How to Examine Using Doppler mode? | 1. Connect the appropriate probe, leaving the probes in their respective holders. 2. Position the patient for the examination. 3. Press Patient. Enter the appropriate patient data using the appropriate exam category and worksheet. 4. Select the preset, application and probe to be used. 5. Locate the anatomy to be examined. Get a good B Mode image. Refer to B-Mode section of this chapter for more information. 6. Press M/D Cursor to display the sample volume cursor and gate. or Press PW. The PW Doppler spectrum appears and the system operates in combined B+Doppler Mode. Adjust Volume to adjust Doppler audio. The Doppler signal is heard through the speakers. 7. Position the sample volume cursor by moving the Trackball left and right. Position the sample volume gate by moving the Trackball up and down. Size the gate by clicking SV Length. 8. Optimize the PW Doppler spectrum, as necessary. Refer to the Doppler Optimization section of this chapter for more information. 9. Press B Pause to toggle between real time B-Mode with sample volume and spectral display with audio. 10. Sample along the whole length of the vessel. Make sure that the probe is parallel to flow. Listen, then look, when positioning the sample volume cursor. 11. Press Freeze to hold the trace in memory and stop imaging. Activate CINE Timeline, as necessary. See ‘Activating CINE’ for more information. 12. Perform measurements and calculations, as necessary. 13. Record results by pressing the appropriate print key, depending on the setup of your recording devices. 14. Press Freeze to resume imaging. 15. Repeat the above procedure until all relevant flow sites have been examined. 16. Replace the probe in its respective holder.
GLOGIQ930 | DOPPLERMODE_ACTIVATION | Activation | optimizing the image | LOGIQ9 | How to Activate Doppler Mode? | To activate PW Doppler Mode, press PW. The Doppler spectrum displays along with the B-Mode image. The cursor changes to a Doppler cursor. You can now position and size the sample volume gate to get a velocity. Use Doppler Audio to listen for when the sample volume gate is positioned over an area of flow.

## Sheet2
TID (unique) | KEYWORD(unique) | TAG(Separate with comma) | CATEGORY | PID(Separate with comma if more than one) | QUESTION | ANSWER | Documentation link | YouTube solution link
GRGXR61 | OPERATOR_CONSOLE | operator console,indicators,indications,radiography,ge | indicator informations | XR6000,pioneer | What does the signals on indicators tell us?. | 1.X-ray ON Indicator Yellow. When system console receives the x-ray on signal, system console will light up the x-ray on indi- cator on the system console, and the x-ray on lamp on the room. In the meantime, system console will turn on the alarm sound. 2.Tube Overheat Indicator Red. When system console receives the tube overheat signal, system console will light this indicator up, and system exposure will be inhibited. When the light is off after appropriate cooling, exposures can be taken. 3.Inhibit Indicator Red. The system inhibit indicator lights up means there is an error. It includes: - Examination room door is not closed - Various inhibition errors on the system - Technique parameters are over the limit 4.Generator Ready Indicator Green. This indicator will light up when system is ready for exposure after pressing prep. If the prep switch is released, the Generator Ready indicator will be off." | https://drive.google.com/file/d/1hROr2w1CToY5uViR9M_WCYJda0AwQJwa/view?usp=sharing
GRGXR62 | WALL_STAND_MODE | Operating mode 1,radiography,ge | modes of operation | XR6000,pioneer | How to apply Wall Stand Mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select wall stand mode . b.) Select focal spot . c.) Adjust exposure parameters. 4. Place Cassette. 5. Adjust wall stand height by patient size. 6. Patient positioning. 7. Adjust tube angulation and SID. 8. Take exposure. 9. Take out cassette, and develop films. | https://drive.google.com/file/d/1FJnqNbkvuWtrrIgBsv1hYsuY6cI6It8R/view?usp=sharing
GRGXR63 | TABLE _MODE | operatig mode 2,radiography,ge | modes of operation | XR6000,pioneer | How to apply Table Mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select table mode . b.) Select focal spot . c.) Adjust exposure parameters.4. Place cassette onto the cassette tray under the table. 5. Patient positioning on the radiographic table.6. Adjust tube angulation and SID.7. Take exposure. 8. Take out cassette, and develop films. | https://drive.google.com/file/d/1JrpHEBzTn4eMrv86Bj7Ww2Kt9cfgznZc/view?usp=sharing
GRGXR64 | TABLE _TOP_MODE | operating mode 3,radiography,ge | modes of operation | XR6000,pioneer | How to apply Table Top Mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select table top mode . b.) Select focal spot . c.) Adjust exposure parameters.4. Patient positioning on radiographic table, and place cassette to the appropriate location. 5. Adjust tube angulation and SID.6. Take exposure. 7. Take out cassette, and develop films. | https://drive.google.com/file/d/1cNFN3-UgLi31TrGLH-qmNStqHJ4-493G/view?usp=sharing
GRGXR65 | STRETCHER_TABLE_MODE | operating mode 4,radiography, ge | modes of operation | XR6000,pioneer | How to apply Stretcher table mode? | 1. System power ON. 2. Tube warm-up. 3. Set technique parameters. a.) Select table top mode . b.) Select focal spot . c.) Adjust exposure parameters. 4. Rotate Tube Stand 180 Degree. 5. Patient positioning on stretcher table, and place cassette to the appropriate location.6. Adjust tube angulation and SI7. Take exposure. 8. Take out cassette, and develop films. | https://drive.google.com/file/d/1xoOaukIM0A1iN67BZuacmmKuG7CuYNjb/view?usp=sharing
GRGXR66 | SYSTEM_POWER_ON | system power options,power on, radiography,ge | sytem power options | XR6000,pioneer | How to power on the system? | Connect the Wall Brake, the POWER ON green indication lamp located on the right side of the POWER ON button on system console lights up. To turn ON the system, press the “POWER ON“ button on the top-right of system console.
GRGXR67 | SYSTEM_POWER_OFF | system power options,power off, radiographyge | sytem power options | XR6000,pioneer | How to power of the system? | To turn OFF the system, press the “POWER OFF“ button on the top-right of system console. All of the subsystems power will be cut off at the same time. System console POWER ON green indication lamp will not be off until the Wall Brake is cut off.
GRGXR68 | SAVING_CHANGES | saving changes,system status,radiography,ge | saving changes | XR6000,pioneer | How to save changes in system parameters? | After change the parameters, press “Save APR“ + “Retrieve APR“ (press “Save APR“ button first) for 3 seconds, system console will save the current parameters as the new system power-on initial mode.
GRGXR69 | TUBE_WARMUP_REQUIREMENTS | tube warm up,requirements,radiography,ge | tube warm up requirements | XR6000,pioneer | What are the tube warm up requirements? | 1. Use the larger focal spot of the tube, and the seasoning will be done with the circuit specified in technical data. 2. The operating tube voltage and current should be monitored all exposures during this seasoning operation. 3. The anode rotation should be utilized at normal speed (3000 or 3600 min-1). 4. Do not exceed either the generator or the tube ratings.
GRGXR610 | TUBE_WARMUP-PROCEDURE | tube warm up, procedures, radiography,ge | tube warm up procedures | XR6000,pioneer | What are the tube warm up procedures? | 1. Provide continuous input of small tube current (Approx.1 mA) and increase the tube voltage stepwise according to the time of the Seasoning Schedule at installation (See Figure 4-1). For any X-ray equipment without the fluoroscopy circuit, perform seasoning of short-time input from 50 kV. 2. Use the larger focus, rotate the anode, set the tube current to 50% of the maximum permissible tube current at maximum tube voltage, apply short-time input once a minute with exposure time of 0.1s, and increase the tube voltage stepwise. When unstable, states of the tube current are observed in the course of increasing tube voltage, reduce the tube voltage to the extent that no variation occurs, and then increase it more slowly than previously. 3. Stain like impact traces might appear on the X-ray tube target surface by the slight discharge in the course of the seasoning. The phenomena are one process to recover the withstanding voltage performance at that time, Even if unstable states appear in the course of the seasoning, the stable operation is obtained at the maxi- Direction 5275602-100 Rev 16 94 mum tube voltage of succeeding seasoning, the tube unit can be used without any interference to its electrical performance. (The maximum tube testing voltage must be strictly limited to the maximum rated tube voltage of the tube unit.) 4. When the X-ray tube unit is used after a dormant period of a week or more, use the tube unit after performing the Seasoning Schedule at Starting . 5. Every day morning, when you start to use X-ray tube, you are requested to warm-up the tube to your necessary high voltage condition for stabilized equipment operation. a.) At least 80kV (Step 1) is requested for all operation. b.) Need voltage up to your necessary daily voltage (Step 2 or higher step). c.) If discharge happened, re-warm up again under the previous condition.
GRGXR611 | CHANGE_APR_PROTOCOL | apr protocol,apr parameters,apr,radiography,ge | change apr protocol | XR6000,pioneer | How to overwrite an existing protocol? | 1. Select the body anatomy button under the APR area on the XR 6000 console. 2. Select any button that needs to be changed (receptor, focus spot size, kVp, mA, mAs, mSec, AEC, exposure view, or patient size). Once a button is selected that is not in the protocol, the body anatomy icon on LCD will flash notifying a change. 3. Press APR Save button for 3 seconds, a beep from system means the new protocol is saved. And the new protocol parameters will be saved as default in system console.
GRGXR612 | TO_MAKE_AN_HANDSWITCH_EXPOSUREN | handswitch,handswitch exposure,radiography,ge | to make an handswitch exposure | XR6000,pioneer | How to make an handswtich exposure? | The handswitch connected to the system console is a three-position push button switch. Its three positions are OFF, Prep and Expose. The handswitch is normally in the OFF position. Press the handswitch halfway to the Prep position for 1~1.5 seconds. This prepares the X-ray tube for exposure. Then press the handswitch all the way down to the Exposure position and hold until the exposure is complete. If release the handswitch during this period, exposure will stop immediately. X-ray on indicator off and beep sound terminated notify that the exposure is complete. Prep on handswitch can be used again. If the operator does not release the Exposure switch after exposure operation, the next exposure will be inhibited. Only after release the Prep switch, a new exposure is permited.
GRGXR613 | TO_MAKE_AN_SYSTEM_CONSOLE_EXPOSURE | system console,system console exposure,radiography,ge | to make an system console exposure | XR6000,pioneer | How to make an system console exposure? | On the lower right hand corner of the system console under Radiography is where the Prep and Exposure buttons are located. Prep: Press this button for 1~1.5 seconds, system will prepare the tube for exposure. Exposure: If the ready indicator lights up, after press this button, the system will generate the X-ray. And system console will light up the x-ray on indicator on the system console and the x-ray on lamp on the room until the exposure is finished. In the meantime, the buzzer will sound until the exposure is finished.
GRGXR614 | TABLE_TOP_MOVEMENT_OPERATING_INSTRUCTIONS | table top,table top movement,operating instructions table top | table top movement operating instructions | XR6000,pioneer | How to move the table top? | The tabletop is locked in place by electromagnetic locks when power is applied. To position the tabletop, the operator must first unlock the electromagnetic locks by pressing the foot pedal located along the base of the patient table two consecutive times . Hold the pedal down releases the longitudinal and lateral locks. To re-lock the tabletop into position, remove your foot from the pedal.
GRGXR615 | CASSETTE_TRAY_MOVEMENT_INSTRUCTIONS | cassette tray,movement instructions | cassette tray movemnt instuctions | XR6000,pioneer | How to move the cassette tray? | The cassette tray can be moved by pressing in the gray lock switch located under the hand grip. Then with the hand grip, move the cassette tray to the desired position. Release the lock switch or the hand grip after the desired position is reached
GRGXR616 | CASSETTE_LOADING | cassette,cassette loading,cassette tray | cassette loading | XR6000,pioneer | How to load a cassette? | To load a cassette, hold the cassette handle and pull out the cassette tray. Place the cassette into the two clamps. Make sure the cassette’s center line and center of cassette tray align.When suitably positioned, lock the cassette by turning the knob clockwise until it is locked. To release and adjust, turn the knob counter clockwise. Finally, push the tray in.
GRGXR617 | CASSETTE_REMOVAL | cassette,cassette tray,cassette removal | cassette removal | XR6000,pioneer | How to remove a cassette? | To remove a cassette from the manual cassette tray, pull the tray out fully by its handle. Turn the knob counter clockwise to release the clamps. Push the cassette in and enlarge the space between the two clamps. The cassette is now free to be removed.
GRGXR618 | X-RAY_TUBE_ASSEMBLY_VERTICAL_MOVEMENT | xray tube,assembly,vertical movement, | xray tube assembly vertical movement | XR6000,pioneer | How vertical movement of xray tube assembly is achieved? | Press in the VERT switch (or ALL) on the Control panel, then hold the handles on the panel and move the tube assembly to the desired position. Notice the mark placed on the tube stand and chose desired SID. After the desired position is reached, release the switch on the tube control panel to re-engage the tube stand vertical locks
GRGXR619 | X-RAY_TUBE_ASSEMBLY_STAND_TRANSVERSE_MOVEMENT | xray tube,assembly,stand transverse movement, | xray tube assembly stand transverse movement | XR6000,pioneer | How transverse movement of xray tube stand is achieved? | Press in the TRANS switch (or ALL) on the Control panel , then hold the handles on the panel and move the tube stand to the desired position. After the desired position is reached, release the switch on the tube control panel, to re-engage the tube stand transverse locks. Remember when the tube stand is moved out of the range of the cassette tray, the cassette tray will decouple from the tube stand
GRGXR620 | X-RAY_TUBE_ASSEMBLY_STAND_ROTATION | xray tube,stand rotation | xray tube asssembly stand rotation | XR6000,pioneer | How to rotate xray tube stand? | The tube stand can rotate to any desired angle +/-180 degrees for off table radiographic images. To rotate the tube stand, grasp the handles on the x-ray tube stand control panel and turn the tube stand to the desired position. The rotation is automatically notched every 90 degrees.
GRGXR621 | X-RAY_TUBE_STAND_ANGULATION | xray tube,angulation | xray tube stand angulation | XR6000,pioneer | How xray tube stand angulation is achieved? | The X-ray tube can angulate +/-120 degrees for angulation views, and there are detents at 0 degree, +/- 90 degrees, and +/-120 degrees. The distance from tube focal spot to the ground is from 555 mm to 1760 mm. A small window on the control panel shows the angulation of the tube. To change the X-ray tube angulation press in and hold the control panel ANG (or ALL) switch and position the tube assembly to the desired angulation. Then release the ANG switch when the desired angulation is reached.
GRGXR622 | X-RAY_TUBE_TILT | xray tube,tilt,movement | xray tube tilt | XR6000,pioneer | How is xray tube tilted? | The X-ray tube can tilt forward more than 30 degrees, and backward more than 20 degrees. Behind the tube are two mechanical locks. Release the locks, the tube is now free to be moved. Position the tube to the desired location and lock the tube in place. To replace the tube to the normal position, release the locks, and position the tube. On the side of the tube, there is an angulation indicator and an arrow to align the tube into correct position.
GRGXR623 | ALLIGNMENT_FIELD_LIGHT | field light,allignment,collimator | allignment field light | XR6000,pioneer | How is field light alligned? | Press the field light button located on the collimator face to activate the high-intensity field lamp.The field light is timed and automatically turns off 30 seconds after your release the switch. To activate the field light for another 30 seconds, press the button again when the light goes out.
GRGXR624 | POSITIONING_OF_COLLIMATOR | postion,collimatore,allignment | positioning of collimatore | XR6000,pioneer | How is a collimator positioned? | Position the collimator and tube using the field light and crosshair shadow as guides. The crosshair shows the field center. Collimate by adjusting the size of the field with the two control knobs located on the front of the collimator.. The left knob controls the transverse dimension of the blades, and the right blade knob controls the longitudinal dimension of the blades. A diagram on the face of the colllimator shows which knob controls which dimension. Field sizes from 0x0 up to 17” x 17” at a 102 cm SID can be obtained. The knob selectors indicate the field size for a selected SID. The maximum field size is 17” x 17”. Dial numbers on the calibrated scale are for 100 cm and 180 cm scale SID. A ruler is located on the left side of the collimator with the measurements in centimeters.
GRGXR625 | COLLIMATOR_ROTATION | rotation,collimator,allignment | collimator rotation | XR6000,pioneer | How is a collimator rotated? | 1. Unlock the Fix Knob located on the top of the collimator and turn the collimator to the desired position. 2. After the collimator is positioned, turn the knob back to the locked position. 3. Unlock the Fix Knob and turn the collimator back to the normal position. Tighten the Fix Knob. Make sure the collimator is in alignment with the cassette tray by turning on the collimator light and seeing that the light is aligned.
GRGXR626 | DAP_METER_OPERATION | dap meter,operation | dap meter operation | XR6000,pioneer | How is a dap meter operated? | By Pressing [Reset] button, all the internal measured value memory is deleted and the display is reset to zero (0.0 or 0.00). The meter is ready for measurement again after [Reset] button is pressed.
GRGXR627 | ADVANCED_WALL_STAND_VERTICAL_POSITIONING | wall stand,allignment,positioning | advanced wall stand vertical positioning | XR6000,pioneer | How are vertical postioning of advanced wall stand achieved? | For advanced wall stand, follow the procedures: 1. Press and hold the switch button to release the lock. 2. Move the wall stand to the desired position. 3. Release the switch button to lock the wall stand.
GRGXR628 | WALL_STAND_CASSETTE_LOADING | wall stand,allignment,positioning,cassette loading | wall stand cassette loading | XR6000,pioneer | How to load a cassette in a wall stand? | 1. Pull the cassette tray out by holding the handle of the cassette tray. 2. Insert the cassette into the tray, and the spring on the cassette tray will automatically center the cassette. 3. Push the cassette tray back into the Tray holder.
GRGXR629 | ADVANCED_WALL_STAND_CASSETTE_LOADING | wall stand casstte tray, cassette loading | advance wall stand casstte loading | XR6000,pioneer | How to load a cassette in an advanced wall stand? | 1. To insert a cassette, pull the tray out of the Bucky to the tray stop. 2. Insert the shelf into the centering scale slots that correspond to the cassette size selected. 3. Lift the clamping lock handle to unlock it. 4. Slide the clamping apart to insert a cassette on the shelf. Clamp and center the cassette transversely on the tray. 5. Push the clamping lock against the cassette and lock it by pressing down the lock handle. 6. Push the cassette tray all the way into the Bucky.
GRGXR630 | STANDARD/ADVANCED_CASSETTE_REMOVAL | wall stand,cassette tray,cassette removal | standard advanced casette removal | XR6000,pioneer | How to remove a cassette ina wall stan? | To remove a cassette from the cassette tray, pull the tray out fully by its handle.